# Calibration Certificate

Customer: -						Report No: BCES/a.o/V.G/0093

M/s. Amneal Oncology Pvt. Ltd.

Jadcherla,Mahaboob Nagar,

Hyderabad- 509302

01.	INSTRUMENT NAME			:	Vacuum Gauge

02.	INSTRUMENT LOCATION			:	Booklet Reader

03.	INSTRUMENT ID				:	EP-PDE-182/VG/03

04.	EQUIPMENT NAME				:	Automatic Cartoning Machine

05.	MAKE						:	Manometer

06.	RANGE					:	0 to -760 mmHg

07.	ACCEPTANCE LIMIT			:	 2% FSD

08.	CALIBRATED ON				:	15-Dec-2021

09.	NEXT DUE DATE				:	14-Dec-2022

10.	ENVIRONMENTAL CONDITIONS		:	Tem: 23.5°C & RH 56.7% (On site)

## Remarks: The instrument is satisfactory with respect to the Specified accuracy for Calibration.

| S. No. | Standard Instrument Value in mmHg | Measured Instrument Value in mmHg | Error in (±) mmHg |
| --- | --- | --- | --- |
| 1 | 0 | 0 | 0 |
| 2 | -99 | -100 | 1 |
| 3 | -198 | -200 | 2 |
| 4 | -297 | -300 | 3 |
| 5 | -595 | -600 | 5 |

Master Instrument used for Calibration: -

Type of Instrument                             : Digital Pressure Gauge

Serial No				: 29651809

Make					: R&D Instruments

Range					: -1 to 40 bar

Traceability                                         : Transcal

Certificate No.				: TSC/21-22/9642 04

Date of Calibration			: 13/Oct/2021

Next Calibration Due On		: 13/Oct/2022

Note: Temperature & Humidity & CO2 Meter: Certificate No.: TC/ 21-22 /936 01

- Tested instrument’s / Standards used are Traceable to National Standards

| Prepared By | Checked By | Reviewed By |
| --- | --- | --- |

Page: 1 of 1

# Calibration Certificate

Customer: -						Report No: BCES/a.o/P.G/3885

M/s. Amneal Oncology Pvt. Ltd.

Jadcherla,Mahaboob Nagar,

Hyderabad- 509302

01.	INSTRUMENT NAME			:	Pressure Gauge

02.	INSTRUMENT LOCATION			:	Glue Gun Machine Side

03.	INSTRUMENT ID				:	EP-PDE-182/PG/03

04.	EQUIPMENT NAME				:	Automatic Cartoning Machine

05.	MAKE						:	Festo

06.	RANGE					:	0 to 16 Bar

07.	ACCEPTANCE LIMIT			:	 2% FSD

08.	CALIBRATED ON				:	15-Dec-2021

09.	NEXT DUE DATE				:	14-Dec-2022

10.	ENVIRONMENTAL CONDITIONS		:	Tem: 23.5°C & RH 56.6% (On site)

## Remarks: The instrument is satisfactory with respect to the Specified accuracy for Calibration.

| S. No. | Standard Instrument Value in Bar | Measured Instrument Value in Bar | Error in (±) Bar |
| --- | --- | --- | --- |
| 1 | 0.000 | 0.0 | 0.000 |
| 2 | 4.005 | 4.0 | 0.005 |
| 3 | 8.009 | 8.0 | 0.009 |
| 4 | 12.013 | 12.0 | 0.013 |
| 5 | 14.018 | 14.0 | 0.018 |

Master Instrument used for Calibration: -

Type of Instrument                             : Digital Pressure Gauge

Serial No				: 29651809

Make					: R&D Instruments

Range					: -1 to 40 bar

Traceability                                         : Transcal

Certificate No.				: TSC/21-22/9642 04

Date of Calibration			: 13/Oct/2021

Next Calibration Due On		: 13/Oct/2022

Note: Temperature & Humidity & CO2 Meter: Certificate No.: TC/ 21-22 /936 01

- Tested instrument’s / Standards used are Traceable to National Standards

| Prepared By | Checked By | Reviewed By |
| --- | --- | --- |

Page: 1 of 1

# Calibration Certificate

Customer: -						Report No: BCES/a.o/P.G/3888

M/s. Amneal Oncology Pvt. Ltd.

Jadcherla,Mahaboob Nagar,

Hyderabad- 509302

01.	INSTRUMENT NAME			:	Pressure Gauge

02.	INSTRUMENT LOCATION			:	Steam Line

03.	INSTRUMENT ID				:	EP-ENE-158/PG/01

04.	EQUIPMENT NAME				:	Pure Steam Generator

05.	MAKE						:	Baumer

06.	RANGE					:	0 to 16 Kg/cm2

07.	ACCEPTANCE LIMIT			:	 1% on FS

08.	CALIBRATED ON				:	17-Dec-2021

09.	NEXT DUE DATE				:	16-Dec-2022

10.	ENVIRONMENTAL CONDITIONS		:	Tem: 23.6°C & RH 56.2% (On site)

## Remarks: The instrument is satisfactory with respect to the Specified accuracy for Calibration.

| S. No. | Standard Instrument Value in Kg/cm2 | Measured Instrument Value in Kg/cm2 | Error in (±) Kg/cm2 |
| --- | --- | --- | --- |
| 1 | 0.000 | 0.0 | 0.000 |
| 2 | 4.005 | 4.0 | 0.005 |
| 3 | 8.013 | 8.0 | 0.013 |
| 4 | 12.019 | 12.0 | 0.019 |
| 5 | 14.024 | 14.0 | 0.024 |

Master Instrument used for Calibration: -

Type of Instrument                             : Digital Pressure Gauge

Serial No				: 29651809

Make					: R&D Instruments

Range					: -1 to 40 bar

Traceability                                         : Transcal

Certificate No.				: TSC/21-22/9642 04

Date of Calibration			: 13/Oct/2021

Next Calibration Due On		: 13/Oct/2022

Note: Temperature & Humidity & CO2 Meter: Certificate No.: TC/ 21-22 /936 01

- Tested instrument’s / Standards used are Traceable to National Standards

| Prepared By | Checked By | Reviewed By |
| --- | --- | --- |

Page: 1 of 1

# Calibration Certificate

Customer: -						Report No: BCES/a.o/P.S/1197

M/s. Amneal Oncology Pvt. Ltd.

Jadcherla,Mahaboob Nagar,

Hyderabad- 509302

01.	INSTRUMENT NAME			:	Pressure Switch

02.	INSTRUMENT LOCATION			:	Feed Pump Outlet Line

03.	INSTRUMENT ID				:	EP-ENE-158/PS/01

04.	EQUIPMENT NAME				:	Pure Steam Generator

05.	MAKE						:           Orion

06.	RANGE					:	0 to 15 Bar

07.	ACCEPTANCE LIMIT			:	 1.5% on FS

08.	CALIBRATED ON				:	17-Dec-2021

09.	NEXT DUE DATE				:	16-Jun-2022

10.	ENVIRONMENTAL CONDITIONS		:	Tem: 24.3°C & RH 56.5% (On site)

## Remarks: The instrument is satisfactory with respect to the Specified accuracy for Calibration.

| S. No. | Standard Instrument Value in Bar | Measured Instrument Value in Bar | Error in (±) Bar |
| --- | --- | --- | --- |
| 1 | 0.000 | 0.0 | 0.000 |
| 2 | 4.004 | 4.0 | 0.004 |
| 3 | 8.009 | 8.0 | 0.009 |
| 4 | 12.013 | 12.0 | 0.013 |
| 5 | 14.018 | 14.0 | 0.018 |

Master Instrument used for Calibration: -

Type of Instrument                             : Digital Pressure Gauge          Digital Multimeter

Serial No				: 29651809                               1604144022

Make					: R&D Instruments                   Waco

Range					: -1 to 40 bar                             Multi

Traceability                                         : Transcal	                                 Transcal

Certificate No.				: TSC/21-22/9642 04               TSC/21-22/9642 03

Date of Calibration			: 13/Oct/2021                           13/Oct/2021

Next Calibration Due On		: 13/Oct/2022                           13/Oct/2022

Note: Temperature & Humidity & CO2 Meter: Certificate No.: TC/ 21-22 /936 01

- Tested instrument’s / Standards used are Traceable to National Standards

| Prepared By | Checked By | Reviewed By |
| --- | --- | --- |

Page: 1 of 1

# Calibration Certificate

Customer: -						Report No: BCES/a.o/C.G/0514

M/s. Amneal Oncology Pvt. Ltd.

Jadcherla,Mahaboob Nagar,

Hyderabad- 509302

01.	INSTRUMENT NAME			:	Compound Gauge

02.	INSTRUMENT LOCATION			:	On PW Storage Tank

03.	INSTRUMENT ID				:	EP-ENE-159/CG/01

04.	EQUIPMENT NAME				:	PW Storage tank and distribution

System (6KL)

05.	MAKE						:	Baumer

06.	RANGE					:	-1 to 6 Kg/cm2

07.	ACCEPTANCE LIMIT			:	 1.6% of F.S

08.	CALIBRATED ON				:	17-Dec-2021

09.	NEXT DUE DATE				:	16-Dec-2022

10.	ENVIRONMENTAL CONDITIONS		:	Tem: 24.6°C & RH 55.9% (On site)

## Remarks: The instrument is satisfactory with respect to the Specified accuracy for Calibration.

| S. No. | Standard Instrument Value in Kg/cm2 | Measured Instrument Value in Kg/cm2 | Error in (±) Kg/cm2 |
| --- | --- | --- | --- |
| 1 | -0.798 | -0.8 | 0.002 |
| 2 | 0.000 | 0.0 | 0.000 |
| 3 | 1.004 | 1.0 | 0.004 |
| 4 | 3.009 | 3.0 | 0.009 |
| 5 | 5.014 | 5.0 | 0.014 |

Master Instrument used for Calibration: -

Type of Instrument                             : Digital Pressure Gauge

Serial No				: 29651809

Make					: R&D Instruments

Range					: -1 to 40 bar

Traceability                                         : Transcal

Certificate No.				: TSC/21-22/9642 04

Date of Calibration			: 13/Oct/2021

Next Calibration Due On		: 13/Oct/2022

Note: Temperature & Humidity & CO2 Meter: Certificate No.: TC/ 21-22 /936 01

- Tested instrument’s / Standards used are Traceable to National Standards

| Prepared By | Checked By | Reviewed By |
| --- | --- | --- |

Page: 1 of 1

# Calibration Certificate

Customer: -						Report No: BCES/a.o/P.T/1605

M/s. Amneal Oncology Pvt. Ltd.

Jadcherla,Mahaboob Nagar,

Hyderabad- 509302

01.	INSTRUMENT NAME			:	Pressure Transmitter

02.	INSTRUMENT LOCATION			:	Pure Steam Line

03.	INSTRUMENT ID				:	EP-ENE-158/PT/01

04.	EQUIPMENT NAME				:	Pure Steam Generator

05.	MAKE						:	Jumo

06.	RANGE					:	-1 to 9 Bar

07.	ACCEPTANCE LIMIT			:	 1% on FS

08.	CALIBRATED ON				:	17-Dec-2021

09.	NEXT DUE DATE				:	16-Jun-2022

10.	ENVIRONMENTAL CONDITIONS		:	Tem: 24.6°C & RH 56.3% (On site)

## Remarks: The instrument is satisfactory with respect to the Specified accuracy for Calibration.

| S. No. | Standard Instrument Value in Bar | Measured Instrument Value in Bar | Error in (±) Bar |
| --- | --- | --- | --- |
| 1 | -0.798 | -0.8 | 0.002 |
| 2 | 0.000 | 0.0 | 0.000 |
| 3 | 2.009 | 2.0 | 0.009 |
| 4 | 4.016 | 4.0 | 0.016 |
| 5 | 8.025 | 8.0 | 0.025 |

Master Instrument used for Calibration: -

Type of Instrument                             : Digital Pressure Gauge

Serial No				: 29651809

Make					: R&D Instruments

Range					: -1 to 40 bar

Traceability                                         : Transcal

Certificate No.				: TSC/21-22/9642 04

Date of Calibration			: 13/Oct/2021

Next Calibration Due On		: 13/Oct/2022

Note: Temperature & Humidity & CO2 Meter: Certificate No.: TC/ 21-22 /936 01

- Tested instrument’s / Standards used are Traceable to National Standards

| Prepared By | Checked By | Reviewed By |
| --- | --- | --- |

Page: 1 of 1

# Calibration Certificate

Customer: -						      Report No: BCES/a.o/UVT/0108

M/s. Amneal Oncology Pvt. Ltd.

Jadcherla, Mahaboob Nagar,

Hyderabad- 509302

01.	INSTRUMENT NAME		            :	UV Transmitter

02.	INSTRUMENT LOCATION			:           Supply Line

03.	INSTRUMENT ID				:	EP-ENE-159/UV/01

04.	EQUIPMENT NAME				:	PW Storage tank and distribution

System (6KL)

05.	RANGE					:	0 to 100 w/m2

06.	MAKE						:	Alfa

07.	ACCEPTANCE LIMIT			:            1% on FS

08.	CALIBRATED ON				:	17-Dec-2021

09.	NEXT DUE DATE				:	16-Jun-2022

10.	ENVIRONMENTAL CONDITIONS            : 	Tem: 24.5°C & RH 56.2% (On site)

## Remarks: The instrument is satisfactory with respect to the Specified accuracy for Calibration.

| S. No. | Standard Instrument Value in Volts | Equivalent Value in w/m2 | Measured Instrument Value in w/m2 | Error in (±) w/m2 |
| --- | --- | --- | --- | --- |
| 1 | 0.00 | 0.0 | 0.0 | 0.0 |
| 2 | 1.00 | 20.0 | 20.0 | 0.0 |
| 3 | 2.00 | 40.0 | 40.1 | 0.1 |
| 4 | 3.00 | 60.0 | 60.1 | 0.1 |
| 5 | 4.00 | 80.0 | 80.2 | 0.2 |

Master Instrument used for Calibration: -

Type of Instrument                        : Multifunction Calibrator

Serial Number			       : 19050343

Make				       : Masibus

Range				       : Multi

Traceability                                    : TransCal

Certificate Number		       : TSC/21-22/10107 02

Date of Calibration		       : 23/Oct/2021

Next Calibration Due On	       : 23/Oct/2022

Note: Temperature & Humidity & CO2 Meter: Certificate No.: TC/ 21-22 /936 01

- Tested equipment’s / Standards used are Traceable to National Standards

| Prepared By | Checked By | Reviewed By |
| --- | --- | --- |

Page: 1 of 1

# Calibration Certificate

Customer: -						      Report No: BCES/a.o/TST/0972

M/s. Amneal Oncology Pvt. Ltd.

Jadcherla, Mahaboob Nagar,

Hyderabad- 509302

01.	INSTRUMENT NAME		            :	Temperature Sensor With

Transmitter

02.	INSTRUMENT LOCATION			:           Boiler Steam Line

03.	INSTRUMENT ID				:	EP-ENE-158/TS/01

04.	EQUIPMENT NAME				:	Pure Steam Generator

05.	RANGE					:	0 to 2000C

06.	MAKE						:	Radix

07.	ACCEPTANCE LIMIT			:            10C

08.	CALIBRATED ON				:	17-Dec-2021

09.	NEXT DUE DATE				:	16-Jun-2022

10.	ENVIRONMENTAL CONDITIONS            : 	Tem: 24.5°C & RH 56.2% (On site)

## Remarks: The instrument is satisfactory with respect to the Specified accuracy for Calibration.

| S. No. | Standard Instrument Value in 0C | Measured Instrument Value in 0C | Error in (±) 0C |
| --- | --- | --- | --- |
| 1 | 25.09 | 25.0 | 0.09 |
| 2 | 40.13 | 40.0 | 0.13 |
| 3 | 80.19 | 80.0 | 0.19 |
| 4 | 120.26 | 120.0 | 0.26 |
| 5 | 160.32 | 160.0 | 0.32 |

Master Instrument used for Calibration: -

Type of Instrument                        : RTD Sensor                                Multifunction Calibrator

Serial Number			       : DIN43760                                   19050343

Make				       : NI	                                     Masibus

Range				       : -70 to 450 0C                               Multi

Traceability                                    : TransCal	                                     TransCal

Certificate Number		       : TSC/21-22/10107 01	             TSC/21-22/10107 02

Date of Calibration		       : 23/Oct/2021                                23/Oct/2021

Next Calibration Due On	       : 23/Oct/2022                                23/Oct/2022

Note: Temperature & Humidity & CO2 Meter: Certificate No.: TC/ 21-22 /936 01

- Tested equipment’s / Standards used are Traceable to National Standards

| Prepared By | Checked By | Reviewed By |
| --- | --- | --- |

Page: 1 of 1

# Calibration Certificate

Customer: -						         Report No: BCES/ a.o /DL/0286

M/s. Amneal Oncology Pvt. Ltd.

Jadcherla, Mahaboob Nagar,

Hyderabad- 509302

01.	INSTRUMENT NAME			             :	T- Type Thermocouple Sensor

02.	CALIBRATION TYPE				:	Pre Calibration

03.	TEMPERATURE RANGE			:	-100 to 2500C

04.	MAKE						:	Kaye

05.	ACCEPTANCE LIMIT				:             0.5°C

06.	CALIBRATED ON				:	18-Dec-2021

07.	ENVIRONMENTAL CONDITIONS	             :            Tem: 24.6°C & RH 56.6% (On site)

Remarks: The equipment is satisfactory with respect to the Specified accuracy for Calibration.

| Serial No. | Standard Instrument Value in 0C | Measured Instrument Value in 0C | Error in (±)0C |
| --- | --- | --- | --- |
| 01 CH: 02 | 90.00 121.10 150.00 | 90.1 121.2 150.2 | 0.10 0.10 0.20 |
| 02 CH: 03 | 90.00 121.10 150.00 | 89.9 121.2 150.2 | 0.10 0.10 0.20 |
| 03 CH: 04 | 90.00 121.10 150.00 | 90.0 121.3 150.1 | 0.00 0.20 0.10 |
| 04 CH: 05 | 90.00 121.10 150.00 | 90.1 121.2 149.8 | 0.10 0.10 0.20 |
| 05 CH: 06 | 90.00 121.10 150.00 | 90.2 121.2 150.2 | 0.20 0.10 0.20 |

Page 1 of 3

Report No: BCES/ a.o /DL/0286

| Serial No. | Standard Instrument Value in 0C | Measured Instrument Value in 0C | Error in (±)0C |
| --- | --- | --- | --- |
| 06 CH: 07 | 90.00 121.10 150.00 | 90.1 121.3 150.2 | 0.10 0.20 0.20 |
| 07 CH: 08 | 90.00 121.10 150.00 | 90.0 121.2 149.8 | 0.00 0.10 0.20 |
| 08 CH: 09 | 90.00 121.10 150.00 | 90.2 121.3 150.1 | 0.20 0.20 0.10 |
| 09 CH: 10 | 90.00 121.10 150.00 | 90.1 121.3 150.3 | 0.10 0.20 0.30 |
| 10 CH: 11 | 90.00 121.10 150.00 | 90.1 121.2 150.1 | 0.10 0.10 0.10 |
| 11 CH: 12 | 90.00 121.10 150.00 | 90.2 121.3 150.2 | 0.20 0.20 0.20 |
| 12 CH: 13 | 90.00 121.10 150.00 | 90.1 121.3 150.1 | 0.10 0.20 0.10 |
| 13 CH: 14 | 90.00 121.10 150.00 | 90.0 121.2 150.2 | 0.00 0.10 0.20 |

Page 2 of 3

Report No: BCES/ a.o /DL/0286

| 14 CH: 15 | 90.00 121.10 150.00 | 90.1 121.2 149.8 | 0.10 0.10 0.20 |
| --- | --- | --- | --- |
| 15 CH: 16 | 90.00 121.10 150.00 | 89.9 121.3 150.2 | 0.10 0.20 0.20 |
| 16 CH: 17 | 90.00 121.10 150.00 | 90.1 121.2 149.8 | 0.10 0.10 0.20 |
| 17 CH: 18 | 90.00 121.10 150.00 | 89.8 121.3 150.1 | 0.20 0.20 0.10 |
| 18 CH: 19 | 90.00 121.10 150.00 | 90.1 121.2 150.2 | 0.10 0.10 0.20 |

Master Instrument used for Calibration: -

Type of Instrument                        : RTD Sensor                                Multifunction Calibrator

Serial Number			       : DIN43760                                   19050343

Make				       : NI	                                     Masibus

Range				       : -70 to 450 0C                               Multi

Traceability                                    : TransCal	                                     TransCal

Certificate Number		       : TSC/21-22/10107 01	             TSC/21-22/10107 02

Date of Calibration		       : 23/Oct/2021                                23/Oct/2021

Next Calibration Due On	       : 23/Oct/2022                                23/Oct/2022

Note: Temperature & Humidity & CO2 Meter: Certificate No.: TC/ 21-22 /936 01

Note:  We are using Hot Bath for Temperature Source

- Tested equipment’s / Standards used are Traceable to National Standards

| Prepared By | Checked By | Reviewed By |
| --- | --- | --- |

Page 3 of 3

# Calibration Certificate

Customer: -						         Report No: BCES/ a.o /DL/0287

M/s. Amneal Oncology Pvt. Ltd.

Jadcherla, Mahaboob Nagar,

Hyderabad- 509302

01.	INSTRUMENT NAME			             :	T- Type Thermocouple Sensor

02.	CALIBRATION TYPE				:	Post Calibration

03.	TEMPERATURE RANGE			:	-100 to 2500C

04.	MAKE						:	Kaye

05.	ACCEPTANCE LIMIT				:             0.5°C

06.	CALIBRATED ON				:	18-Dec-2021

07.	ENVIRONMENTAL CONDITIONS	             :            Tem: 24.3°C & RH 56.5% (On site)

Remarks: The equipment is satisfactory with respect to the Specified accuracy for Calibration.

| Serial No. | Standard Instrument Value in 0C | Measured Instrument Value in 0C | Error in (±)0C |
| --- | --- | --- | --- |
| 01 CH: 02 | 30.00 60.00 90.00 105.00 | 30.1 60.2 90.1 105.2 | 0.10 0.20 0.10 0.20 |
| 02 CH: 03 | 30.00 60.00 90.00 105.00 | 30.2 60.1 90.1 105.2 | 0.20 0.10 0.10 0.20 |
| 03 CH: 04 | 30.00 60.00 90.00 105.00 | 30.1 60.1 90.2 105.1 | 0.10 0.10 0.20 0.10 |
| 04 CH: 05 | 30.00 60.00 90.00 105.00 | 30.2 60.1 90.2 105.2 | 0.20 0.10 0.20 0.20 |

Page 1 of 3

Report No: BCES/ a.o /DL/0287

| Serial No. | Standard Instrument Value in 0C | Measured Instrument Value in 0C | Error in (±)0C |
| --- | --- | --- | --- |
| 05 CH: 06 | 30.00 60.00 90.00 105.00 | 30.1 60.1 90.2 105.1 | 0.10 0.10 0.20 0.10 |
| 06 CH: 07 | 30.00 60.00 90.00 105.00 | 29.9 60.1 90.2 105.2 | 0.10 0.10 0.20 0.20 |
| 07 CH: 08 | 30.00 60.00 90.00 105.00 | 29.8 60.1 90.1 105.2 | 0.20 0.10 0.10 0.20 |
| 08 CH: 09 | 30.00 60.00 90.00 105.00 | 30.0 60.1 90.1 105.2 | 0.00 0.10 0.10 0.20 |
| 09 CH: 10 | 30.00 60.00 90.00 105.00 | 30.1 60.2 90.1 105.2 | 0.10 0.20 0.10 0.20 |
| 10 CH: 11 | 30.00 60.00 90.00 105.00 | 30.2 60.2 90.1 105.3 | 0.20 0.20 0.10 0.30 |

Page 2 of 3

Report No: BCES/ a.o /DL/0287

| Serial No. | Standard Instrument Value in 0C | Measured Instrument Value in 0C | Error in (±)0C |
| --- | --- | --- | --- |
| 11 CH: 12 | 30.00 60.00 90.00 105.00 | 30.1 60.1 90.0 105.1 | 0.10 0.10 0.00 0.10 |
| 12 CH: 13 | 30.00 60.00 90.00 105.00 | 30.2 60.1 89.9 105.2 | 0.20 0.10 0.10 0.20 |
| 13 CH: 14 | 30.00 60.00 90.00 105.00 | 30.1 60.1 90.2 105.2 | 0.10 0.10 0.20 0.20 |

Master Instrument used for Calibration: -

Type of Instrument                        : RTD Sensor                                Multifunction Calibrator

Serial Number			       : DIN43760                                   19050343

Make				       : NI	                                     Masibus

Range				       : -70 to 450 0C                               Multi

Traceability                                    : TransCal	                                     TransCal

Certificate Number		       : TSC/21-22/10107 01	             TSC/21-22/10107 02

Date of Calibration		       : 23/Oct/2021                                23/Oct/2021

Next Calibration Due On	       : 23/Oct/2022                                23/Oct/2022

Note: Temperature & Humidity & CO2 Meter: Certificate No.: TC/ 21-22 /936 01

Note:  We are using Hot Bath for Temperature Source

- Tested equipment’s / Standards used are Traceable to National Standards

| Prepared By | Checked By | Reviewed By |
| --- | --- | --- |

Page 3 of 3